EVAL-CN0300-EB1Z TEST PROCEDURE

- If you are running this test procedure for the first time on your PC you need to ensure your PC has the USB drivers for the USB-SWD/UART board.

  - To Install the drivers for the USB-SWD/UART board, unzip the CDM 2.04.06~.zip file for Windows XP drivers

  - For other operating systems, go to the FTDI website:

    - http://www.ftdichip.com/Drivers/D2XX.htm

- Connect the EVAL-CN0300-EB1Z board to the USB-SWD/UART board in figure 1.

- Connect a 12V power supply to J3. Make sure to connect a Current Meter in series with the 12V terminal. Short the inputs to J2 together.

4.  Connect the USB-SWD/UART board to your PC via a USB cable. (Ensure the driver is installed correctly)

5.  You now need to determine what com port your PC has allocated to your USB-SWD/UART board

a) Go to “Settings, Control panel, System, Hardware, Device Manager – see example below where COM port 56 was allocated.

6.  Next open the application CM3WSD.exe

7.  Using the “Browse” tab, call the file CN0300_test.hex

8.  Select the “Configure” tab and

A) select the ADuCM360 in “Parts” tab.

B) Select “Comms” tab. Select the appropriate COM port as determined in step 5. A baud rate of 9600 is fine.

C) Select “Command” tab. Enable the “program and verify” tabs.

9.  Next, force the part into download mode – to do this, hold the “SD” button on the CN0300-EVAL-EB1Z board while pressing and releasing the “RESET” button.

10. Click the “Start” button on CM3WSD. If the download has started successfully, you will see the following:

If this doesn’t happen, repeat step 9.

11. Next close CM3WSD and open HypterTerminal on your PC.

Set Hyperterminal for the appropriate COM port and for 9600 baud rate, 8 data bits, No parity, 1 stop bit, no flow control

12. Press reset and the following should appear on your PC:

13. Your current meter should report a value close to 4mA – if it doesn’t follow the instructions on the HyperTerminal screen to calibrate the DAC output to 4mA by pressing the “1” or “0” keys accordingly. Ensure you select the Hyperterminal screen before pressing 1 or 0.

Press return when the current meter reports 4mA

14. Repeat step 13 for the 20mA calibration.

15. After this the program will continuously output the following information:

The Final temperature reported should be within a few degrees C of the expected ambient temperature. +/-5C is fine since the ADC is not fully calibrated.

16. Temporarily disconnect the COM port from HyperTerminal by pressing the phone tab.

17. Repeat Steps 7 to 10 except this time, select the file CN0300_Final.hex. This file is very similar to the test program except it does not call the DAC calibration function.

18. Close CM3WSD. Return to HyperTerminal and select “Connect” by pressing the phone icon.

19. Press the reset button on the EVAL-CN0300-EB1Z board and the following should appear on the HyperTerminal window: the current reading on the current meter should match the value “Expected DAC Current” on your screen.