Document No.	:	  18-051370-01 Rev A

Title	:	HMC8414-EVALZ Customer Evaluation Board Test Procedure

| REVISION HISTORY |  |  |  |  |
| --- | --- | --- | --- | --- |
| Revision | ECR # | Description of Change | Date | Author |
| A | ECR-119098 | For Release to Manufacturing. ECN will be approved by: Najie Pandian | 06Feb24 |  |

| Required Approvers |  |
| --- | --- |
| Approver Roles | Approver Names |
| Product Engineer |  |

| Model Name | C-Spec | Tester Program |
| --- | --- | --- |
| HMC8414-EVALZ | C-05949 D | HMC8414-EVALZ |

Test 1: S-Parameter Sweep Equipment Needed:

3 Power Supply – Keithley 2400 or its equivalent

1 2-port Network Analyzer – N5245 or its equivalent

# Setup:

Connect the Network Analyzer as shown below:

Port 1 – RFIN Port 2 – RFOUT

# Network Analyzer Settings:

Min/Max Frequency: 100 MHz to 10 GHz Number of Points: 401 points

Input Power: -20 dBm IF BW: 1kHz

# Setting up notes:

- Calibrate the Network Analyzer to compensate for the loss of both input and output cables.

- Record the S11, S12, S21(Gain), S22 of the DUT. Compare the data with the plot below.

# Biasing steps:

For the LNA path:

- Turn on VBIAS = + 5V, typical IBIAS = ~3 - 5 mA (SMU Compliance at 7.5 mA)

- Turn on VDD = + 5V, typical IDD = ~ 70 – 110 mA (SMU Compliance at 120 mA)

- Turn on VCTRL = + 5V, typical ICTRL = ~ 0.4 mA (SMU Compliance at 2 mA)

- Apply the RF signal.

- Log the data.

- Turn off the RF signal.

# For the Bypass path:

- Retain the value of VBIAS and VDD and change the value of VCTRL = 0V, typical ICTRL = ~0 mA, IBIAS = ~0 mA, IDD = ~3 mA.

- Apply the RF signal.

- Log the data.

# Power off:

- Turn off the RF signal.

- Turn off VCTRL.

- Turn off VDD.

- Turn off VBIAS.

# S21, S11 and S22 Plot

Test 2: ATE Setup

# Equipment and Accessories Needed:

3 Power Supply – Keithley 2400 or its equivalent

1 Signal Generator – Keysight E8267D or its equivalent

1 Power meter – Anritsu ML2496A or its equivalent

2 Power Sensors – Anritsu MA2445D or its equivalent 2 DMM – Agilent 34401A or its equivalent

2 Bias Tee – Marki BTN20050 or its equivalent 1 10 dB pad

1 Coupler – Krytar or equivalent

1 external amplifier – Keysight N4985a or equivalent Connect the setup as shown below.

Test Parameters: