Document No.	:	18-080958-01 Rev A

Title			:	ADL8122-EVAL1Z Customer Evaluation Board Test Procedure

| REVISION HISTORY |  |  |  |  |
| --- | --- | --- | --- | --- |
| Revision | ECR # | Description of Change | Date | Author |
| A |  | Initial Release | 04/25/24 | Joe Russell Romero |

| Required Approvers |  |
| --- | --- |
| Approver Roles | Approver Names |
| Product Engineer | Joe Russell Romero |

Test 1: S-parameter Sweep:

Equipment Needed:

1 Power Supply: Keithley 2400 or equivalent – PS1

1x 2-port Network Analyzer – N5245B or its equivalent

Setup:

Connect the Network Analyzer as shown;

Port 1 – RFin

Port 2 – RFout

Connect the Keithley Power supplies as shown:

+ output: VDD

Ground: GND

Network Analyzer Settings:

Min/Max Frequency: 10MHz – 14GHz

Number of Points: 401 points

Input Power: -25dBm

IF Bandwidth: 1kHz

Setting up notes:

- Calibrate the Network Analyzer to compensate for the loss of both input and output cables.

- Record the S11, S12, S21(Gain), S22 of the DUT.

- Measure the S21 of the thru cal path and compensate the loss of the cal path in the measurement of the DUT.

- Compare the data recorded for S21 across evaluation boards to the reference S21 plot below.

There should be no dips nor spikes and it should be within the limits (refer to the next test section).

Bias the part according to the Biasing steps below. 
Biasing steps:

Power on:

- Set VDD = 5V

- Measure and record Itotal, typical of 95mA (LL: 80mA, UL:110mA)

- Apply the RF signal.

Power off:

- Turn off the RF signal.

- Set VDD = 0V

Reference Plots

PNA-X:

ATE:

Test 2: ATE Test Setup:

| Model Name | C-Spec | Tester Program |
| --- | --- | --- |
| ADL8122-EVAL1Z | *Ongoing Creation* | ADL8122-EVAL1Z.par |

Test set-up:

Test Parameters: